# ASSESSMENT

Data Integrity Risk Assessment will be completed for equipment/systems to determine the level of risk to generated data. The assessment shall be completed by System Owners or designee with support from system SMEs and Quality Assurance (Data Integrity Officer / Manager), as appropriate.

Data Integrity Risk Assessments are model and software specific. Systems that are identical, may share the same risk assessment if it can be demonstrated that they have the same configuration and process controls.

Complete the following tables for the intended system/equipment:

| Table 1 - System Details |  |
| --- | --- |
| Equipment /System Type: |  |
| Manufacturer: |  |
| Model: |  |
| Equipment /System Name: |  |
| Asset ID (main system): |  |
| Owner Department: |  |
| System Owner: |  |
| Software Application Name: |  |
| Software Application Version: |  |
| Operating System Detail (e.g., Windows 7, Windows 10): |  |
| Host Name (Computer Name): |  |
| IP Address (if static) and Network: |  |
| List the types of data recorded and stored by the system: |  |

# Assessment Criteria

The risk assessment will focus on the following 10 categories:

| - Operating System Security | - Data Folder Security |
| --- | --- |
| - Software Security | - User Level and Privileges Control Criteria |
| - Data Backup Criteria | - Password Control Criteria |
| - Audit Trail Criteria | - Methods (or Process Control Parameters) Control Criteria |
| - Control of Date and Time | - System Governance Procedures |

Each category is comprised of a series of control elements that are weighted according to their severity (criticality) from a data integrity compliance perspective. The assessment is designed to ask simple yes/no or true/false type questions of the system. For each control element that scores 5 (very high severity), a risk-based justification or a mitigation action is required. Mitigation actions should focus on reducing the occurrence of the control element and/or improving the detectability. Overall, these will reduce potential patient impact.

It is the responsibility of the Process and System Owner to implement the mitigation actions.

| Table 1 – Risk Rating System |  |  |  |  |
| --- | --- | --- | --- | --- |
| Risk Rating | Severity | Control Element | Patient (product quality) Impact | Detectability |
| 5 | Very High | As defined through yes/no answers a relative weighting for each control element. | Patients are endangered, potential for product recall | The failure is unlikely to be detected |
| 4 | High |  | Critical deviation must be reported in the appropriate system. | The failure can only be found during deep dive assessments |
| 3 | Moderate |  | Non-critical deviation still reported in the appropriate system | The failure may be detected through routine audits |
| 2 | Low |  | Technical problem-Reported as DI event | The failure can be detected through self-audits or data/batch record review. |
| 1 | Very Low |  | No impact on product quality | An automated system is in place that monitors and alerts for failures |

Additionally, each system will be assed for compliance with the following elements:

- Review of Electronic Data and audit trails

- Secure backup of all data and meta data

- Controlled access to computerized systems (segregation of duties)

- Standardization of time sources

- Secure date and time

:

# Operating System Security

| Table 2- Operating System Security Risk Assessment |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Item | Control Element |  | Meets Requirement (Yes / No or N/A) | Score | Outline mitigation strategy or risk-based justification for any control element that scores 5, or comment for clarification purposes. |
|  | Does the system have a computer Operating System (OS) such as Win7 or higher (currently supported by Manufacturer)? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |
|  | Does every user have unique login credentials, if applicable? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |
|  | Is the system connected to the Active Directory Network? |  | ☐ Yes: Score = 1 ☐ No: Score = 3 ☐ N/A: Score = N/A |  |  |
|  | Are all USB ports disabled? |  | ☐ Yes: Score = 1 ☐ No: Score = 3 ☐ N/A: Score = N/A |  |  |
|  | Is the configuration of the Operating System defined and documented in the validation documentation? (e.g., computer name, Active Directory groups, policies, local user access, desktop security, organization unit) |  | ☐ Yes: Score = 1 ☐ No: Score = 3 ☐ N/A: Score = N/A |  |  |
|  | Are Operating System updates restricted by design and configuration of the system? For example, by Active Directory organizational units or by segregation from network connectivity |  | ☐ Yes: Score = 1 ☐ No: Score = 3 ☐ N/A: Score = N/A |  |  |
|  | Is local system administrator access (OS) limited to ONLY IT personnel? |  | ☐ Yes: Score = 1 ☐ No: Score = 3 ☐ N/A: Score = N/A |  |  |
|  | If No, does the local administrator (OS) report outside the department generating the data? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |
|  | Is the OS access configured to ensure that every user uses a unique and attributable account? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |
|  | Are generic local OS Admin accounts limited to appropriate IT? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |
|  | Are the access permissions to the OS controlled by procedure and validated? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |
|  | Is the OS desktop locked (secured) from access to ‘normal’ users? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |
|  | Are Windows hotkeys (i.e., keypad shortcuts) disabled for ‘normal‘ users? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |
|  | Is Windows Explorer disabled for ‘normal‘ users? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |
|  | Is Windows registry editor disabled for ‘normal’ users?? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |
|  | Is access to change the OS date/time disabled to ‘normal’ users? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |
|  | Is access to change the time zone disabled to ‘normal’ users? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |
|  | Is the ability to copy, delete, or rename files disabled for all users? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |
|  | Is there any third-party software installed capable of manipulating images or data (e.g., paint, publisher, Microsoft office) available to the user? |  | ☐ Yes: Score = 5 ☐ No: Score = 1 ☐ N/A: Score = N/A |  |  |
|  | Is antivirus protection installed and active? |  | ☐ Yes: Score = 1 ☐ No: Score = 3 ☐ N/A: Score = N/A |  |  |
|  | Does antivirus protection include automatic updates? |  | ☐ Yes: Score = 1 ☐ No: Score = 3 ☐ N/A: Score = N/A |  |  |
| Average Control Element Score: |  |  |  |  | Report average score, rounded up/down to the nearest whole number. |
| Patient Impact (overall): |  |  |  |  | What is the impact of any control element with a score of 3 or greater? |
| Detectability (overall): |  |  |  |  | What is the detectability of any control element with a score of 3 or greater? |
| Overall Risk Rating = Average Control Element Score x Patient Impact x Detectability Score: |  |  |  |  |  |
| Additional Comments: |  |  |  |  |  |
| Table completed by (Initial and Date): |  |  |  |  |  |

# Software System Security Risk

| Table 3 – Software System Security Risk Assessment and Rating |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Item | Control Element |  | Meets Requirement (Yes / No or N/A) |  | Score | Outline mitigation strategy or risk-based justification for any control element that scores 5 or comment for clarification purposes. |
|  | Does every user have unique login credentials for the application including all administrators? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
|  | Are shared (non-attributable) accounts used during validation disabled? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
|  | Are there any shared/generic logins used by the vendor? |  | ☐ Yes: Score = 5 ☐ No: Score = 1 ☐ N/A: Score = N/A |  |  |  |
|  | Is the shared/generic vendor login and password known by users? |  | ☐ Yes: Score = 5 ☐ No: Score = 1 ☐ N/A: Score = N/A |  |  |  |
|  | Is vendor access managed via SOP, so that it is possible to link actions to a single person? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
|  | Are the user permission settings that allow ’Hidden fields’, ‘Voiding of data’, ‘Overwriting data’, ‘deletion of data’ disabled or controlled within the system? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
|  | Are ‘annotation tools’ disabled in the application to prevent obscuring data? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
|  | Can data be compromised (i.e., deleted, changed, or modified) within the system during normal operation? |  | ☐ Yes: Score = 5 ☐ No: Score = 1 ☐ N/A: Score = N/A |  |  |  |
|  | Is data acquired electronically? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
|  | Is data acquired approved by the user? |  | ☐ Yes: Score = 1 ☐ No: Score = 3 ☐ N/A: Score = N/A |  |  |  |
|  | Is the data saved automatically? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
|  | Is data preview mode available? |  | ☐ Yes: Score = 5 ☐ No: Score = 1 ☐ N/A: Score = N/A |  |  |  |
|  | Do users have access to the data folder outside of the software, for example, using Windows Explorer, or command line prompt? |  | ☐ Yes: Score = 5 ☐ No: Score = 1 ☐ N/A: Score = N/A |  |  |  |
|  | Are electronic signatures available? |  | ☐ Yes: Score = 1 ☐ No: Score = 1 ☐ N/A: Score = N/A |  |  |  |
|  | If so, are they being used? |  | ☐ Yes: Score = 1 ☐ No: Score = 3 ☐ N/A: Score = N/A |  |  |  |
|  | Are biometrics available? |  | ☐ Yes: Score = 1 ☐ N/A: Score = N/A |  |  |  |
|  | If so, are they being used with double authentication (e.g., fingerprint and heartrate)? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
| Average Control Element Score: |  |  |  |  |  | Report average score, rounded up/down to the nearest whole number. |
| Patient Impact (overall): |  |  |  |  |  | What is the impact of any control element with a score of 3 or greater? |
| Detectability (overall): |  |  |  |  |  | What is the detectability of any control element with a score of 3 or greater? |
| Overall Risk Rating = Average Control Element Score x Patient Impact x Detectability Score: |  |  |  |  |  |  |
| Additional Comments: |  |  |  |  |  |  |
| Table completed by (Initial and Date): |  |  |  |  |  |  |

# Data Backup Criteria Risk

| Table 4 – Data Backup Criteria Risk Assessment and Rating |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Item | Control Element |  | Meets Requirement (Yes / No or N/A) |  | Score | Outline mitigation strategy or risk-based justification for any control element that scores 5 or comment for clarification purposes. |
| 1 | Is the backup process outlined in a procedure (SOP)? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
| 2 | Has the data retention time period been defined? |  | ☐ Yes: Score = 1 ☐ No: Score = 3 ☐ N/A: Score = N/A |  |  |  |
| 3 | Is the backup is done automatically by the system? |  | ☐ Yes: Score = 1 ☐ No: Score = 3 ☐ N/A: Score = N/A |  |  |  |
| 4 | Does the backup include all data? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
| 5 | Does the backup include all metadata? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
| 6 | Does the backup include all audit trails (e.g., event & alarm logs)? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
| 7 | Is the backup is performed by IT? |  | ☐ Yes: Score = 1 ☐ No: Score = 3 ☐ N/A: Score = N/A |  |  |  |
| 8 | Is the backup stored on an IT server? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
| 9 | Does the system have an SOP governing archival and retention of data? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
| 10 | Does the system have a defined archive process? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
| 11 | What is the frequency at which Backups occur (answer one only)? |  |  |  |  |  |
|  | - Daily (or less) |  | ☐ Yes: Score = 1 ☐ No: Score = N/A |  |  |  |
|  | - Weekly |  | ☐ Yes: Score = 2 ☐ No: Score = N/A |  |  |  |
|  | - Monthly |  | ☐ Yes: Score = 3 ☐ No: Score = N/A |  |  |  |
|  | - Quarterly |  | ☐ Yes: Score = 4 ☐ No: Score = N/A |  |  |  |
|  | - Annually or greater; or no defined schedule |  | ☐ Yes: Score = 5 |  |  |  |
| Average Control Element Score: |  |  |  |  |  | Report average score, rounded up/down to the nearest whole number. |
| Patient Impact (overall): |  |  |  |  |  | What is the impact of any control element with a score of 3 or greater? |
| Detectability (overall): |  |  |  |  |  | What is the detectability of any control element with a score of 3 or greater? |
| Overall Risk Rating = Average Control Element Score x Patient Impact x Detectability Score: |  |  |  |  |  |  |
| Additional Comments: |  |  |  |  |  |  |
| Table completed by (Initial and Date): |  |  |  |  |  |  |

# Audit Trail Criteria Risk

| Table 5 – Audit Trail Criteria Risk Assessment and Rating |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Item | Control Element |  | Meets Requirement (Yes / No or N/A) |  | Score | Outline mitigation strategy or risk-based justification for any control element that scores 5 or comment for clarification purposes. |
|  | Does the system have an electronic audit trail? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
|  | Is the audit trail enabled? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
|  | Can the audit trail be turned off only by administrators? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
|  | Does the audit trail capture all events (e.g., changes, rename, delete, errors)? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
|  | Does the audit trail capture the identification of who made the change and when the change was made (i.e., date and time)? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
|  | Does the audit trail show the reason (comment) why the change was made? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
|  | Is the audit trail viewable by all users? |  | ☐ Yes: Score = 1 ☐ No: Score = 3 ☐ N/A: Score = N/A |  |  |  |
|  | Is the audit trail protected from changes by all users? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
|  | Is the Windows event log used by the software to support audit trails? |  | ☐ Yes: Score = 5 ☐ No: Score = 1 ☐ N/A: Score = N/A |  |  |  |
|  | If the Windows event log is used, is the use of the event log described in an SOP and validated? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
|  | If the Windows Event Log is used, is the configuration of the event log included in the documented configuration of the computerized system? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
|  | If the Windows Event log is used, is the event log set to be overwritten? |  | ☐ Yes: Score = 5 ☐ No: Score = 1 ☐ N/A: Score = N/A |  |  |  |
|  | If the Windows Event log is used, is it backed-up and archived? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
|  | If used, is the Windows Event Log assessed and reviewed? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
| Average Control Element Score: |  |  |  |  |  | Report average score, rounded up/down to the nearest whole number. |
| Patient Impact (overall): |  |  |  |  |  | What is the impact of any control element with a score of 3 or greater? |
| Detectability (overall): |  |  |  |  |  | What is the detectability of any control element with a score of 3 or greater? |
| Overall Risk Rating = Average Control Element Score x Patient Impact x Detectability Score: |  |  |  |  |  |  |
| Additional Comments: |  |  |  |  |  |  |
| Table completed by (Initial and Date): |  |  |  |  |  |  |

# Data Folder Security Risk

| Table 6 – Data Folder Security Risk Assessment and Rating |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Item |  | Control Element |  |  | Meets Requirement (Yes / No or N/A) |  | Score | Outline mitigation strategy or risk-based justification for any control element that scores 5 or comment for clarification purposes. |
|  | Is the data generated by the system stored in a relational database (SQL/Oracle)? |  |  | ☐ Yes: Score = 1 ☐ No: Score = 3 ☐ N/A: Score = N/A |  |  |  |  |
|  | Is the data generated by the system stored in a file folder structure commonly referred to as flat files? |  |  | ☐ Yes: Score = 3 ☐ No: Score = 1 ☐ N/A: Score = N/A |  |  |  |  |
|  | Where is the data stored? - Is the data stored on the local computer? |  |  | ☐ Yes: Score = 5 ☐ No: Score = 1 ☐ N/A: Score = N/A |  |  |  |  |
|  | - Is the data stored on a network server maintained by IT? |  |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |  |
|  | Can the data folders be ‘copied’, ‘deleted’, ‘saved as’ or moved to other folders by operators/users? |  |  | ☐ Yes: Score = 5 ☐ No: Score = 1 ☐ N/A: Score = N/A |  |  |  |  |
|  | Are Operating System commands such as ‘save’, ‘save as’, ‘copy’, ‘delete’, recycle bin etc., options disabled to operators/users? |  |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |  |
|  | Data can be reviewed and signed off electronically |  |  | ☐ Yes: Score = 1 ☐ No: Score = 1 ☐ N/A: Score = N/A |  |  |  |  |
|  | Can data and metadata (e.g., audit trails) be printed? |  |  | ☐ Yes: Score = 1 ☐ No: Score = 3 ☐ N/A: Score = N/A |  |  |  |  |
|  | Is this an open or closed system? |  |  | ☐ Closed:Score = 1 ☐ Open: Score = 3 ☐ N/A: Score = N/A |  |  |  |  |
|  | If open, is data encrypted? |  |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |  |
|  | If open, are digital signatures used? |  |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |  |
|  | Do signed documents contain the full printed name, date and time, and meaning of the signature? |  |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |  |
|  | If not, is this information traceable within the system (i.e., audit trail)? |  |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |  |
|  | If traceable, is the electronic record in a human readable format? |  |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |  |
|  | Are electronic signatures unique to an individual? |  |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |  |
| Average Control Element Score: |  |  |  |  |  |  |  | Report average score, rounded up/down to the nearest whole number. |
| Patient Impact (overall): |  |  |  |  |  |  |  | What is the impact of any control element with a score of 3 or greater? |
| Detectability (overall): |  |  |  |  |  |  |  | What is the detectability of any control element with a score of 3 or greater? |
| Overall Risk Rating = Average Control Element Score x Patient Impact x Detectability Score: |  |  |  |  |  |  |  |  |
| Additional Comments: |  |  |  |  |  |  |  |  |
| Table completed by (Initial and Date): |  |  |  |  |  |  |  |  |

# Control of Date and Time

| Table 7 – Control of Date and Time Risk Assessment and Rating |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Item | Control Element |  | Meets Requirement (Yes / No or N/A) |  | Score | Outline mitigation strategy or risk-based justification for any control element that scores 5 or comment for clarification purposes. |
| 1 | Is the Operating System synchronized to the network time service? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
| 2 | If No, does a procedure exist to verify the date and time periodically (e.g., prior to use or daylight savings time)? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
| 3 | Access to change the date/time is not available to everyone but those with “Admin” control. |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
| 4 | Access to change the time zone is not available to everyone but those with “Admin” control. |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
| Average Control Element Score: |  |  |  |  |  | Report average score, rounded up/down to the nearest whole number. |
| Patient Impact (overall): |  |  |  |  |  | What is the impact of any control element with a score of 3 or greater? |
| Detectability (overall): |  |  |  |  |  | What is the detectability of any control element with a score of 3 or greater? |
| Overall Risk Rating = Average Control Element Score x Patient Impact x Detectability Score: |  |  |  |  |  |  |
| Additional Comments: |  |  |  |  |  |  |
| Table completed by (Initial and Date): |  |  |  |  |  |  |

# User Level and Privileges Control Criteria

| Table 8 – User Level and Privileges Control Criteria Risk Assessment and Rating |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Item | Control Element |  | Meets Requirement (Yes / No or N/A) |  | Score | Outline mitigation strategy or risk-based justification for any control element that scores 5 or comment for clarification purposes. |
| 1 | Tiered User Management (select 1 below) |  |  |  |  |  |
|  | - System supports only 1 User Type. |  | ☐ Yes: Score = 5 ☐ No: Score = N/A |  |  |  |
|  | - System supports 2 User Types. |  | ☐ Yes: Score = 4 ☐ No: Score = N/A |  |  |  |
|  | - System supports 3 User Types. |  | ☐ Yes: Score = 3 ☐ No: Score = N/A |  |  |  |
|  | - System supports 4 or more User Types. |  | ☐ Yes: Score = 1 ☐ No: Score = N/A |  |  |  |
| 2 | System does not support multi-user type assignment to a single user. |  | ☐ Yes: Score = 5 ☐ No: Score = N/A |  |  |  |
| 3 | Do Administrators report to the department that is responsible for generating/acquiring the data? |  | ☐ Yes: Score = 5 ☐ No: Score = 1 ☐ N/A: Score = N/A |  |  |  |
| 4 | Does the administrator have the ability to generate GMP data? |  | ☐ Yes: Score = 3 ☐ No: Score = 1 ☐ N/A: Score = N/A |  |  |  |
| 5 | Are users appropriately trained for their access level? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
| Average Control Element Score: |  |  |  |  |  | Report average score, rounded up/down to the nearest whole number. |
| Patient Impact (overall): |  |  |  |  |  | What is the impact of any control element with a score of 3 or greater? |
| Detectability (overall): |  |  |  |  |  | What is the detectability of any control element with a score of 3 or greater? |
| Overall Risk Rating = Average Control Element Score x Patient Impact x Detectability Score: |  |  |  |  |  |  |
| Additional Comments: |  |  |  |  |  |  |
| Table completed by (Initial and Date): |  |  |  |  |  |  |

# Password Control Criteria Risk

| Table 9 – Password Control Criteria Risk Assessment and Rating |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Item | Control Element |  | Meets Requirement (Yes / No or N/A) |  | Score | Outline mitigation strategy or risk-based justification for any control element that scores 5 or comment for clarification purposes. |
|  | Does the system use the Active Directory (LDAP) for user credentials and authentication? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
|  | Are user access accounts and credentials managed locally (i.e., not using Active Directory)? |  | ☐ Yes: Score = 3 ☐ No: Score = 1 ☐ N/A: Score = N/A |  |  |  |
|  | If user access credentials are managed (locally), do they meet the following minimum requirements? Use comments section to current outline configuration and if the application has the ability to meet requirements. - Is the system capable of 8 characters minimum? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
|  | - Can the system have complexity requirements (e.g., uppercase, lowercase, numbers, special characters)? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
|  | - Is the system capable of setting up password to be changed every 90 days? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
|  | - Will the system be locked out after 5 unsuccessful attempts? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
|  | - Does the system prevent user to reuse last 5 passwords? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
| Average Control Element Score: |  |  |  |  |  | Report average score, rounded up/down to the nearest whole number. |
| Patient Impact (overall): |  |  |  |  |  | What is the impact of any control element with a score of 3 or greater? |
| Detectability (overall): |  |  |  |  |  | What is the detectability of any control element with a score of 3 or greater? |
| Overall Risk Rating = Average Control Element Score x Patient Impact x Detectability Score: |  |  |  |  |  |  |
| Additional Comments: |  |  |  |  |  |  |
| Table completed by (Initial and Date): |  |  |  |  |  |  |

# Method (or process controls parameter) Control Criteria

| Table 10 – Method (or process controls parameter) Control Criteria Risk Assessment and Rating |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Item | Control Element | Meets Requirement (Yes / No or N/A) |  |  | Score | Outline mitigation strategy or risk-based justification for any control element that scores 5 or comment for clarification purposes. |
|  | Can ‘normal” users / operators modify or overwrite saved recipes/methods/protocols? | ☐ Yes: Score = 5 ☐ No: Score = 1 ☐ N/A: Score = N/A |  |  |  |  |
|  | Are obsolete or superseded recipes/methods/protocols locked from use by a ‘normal’ user/operator? | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |  |
|  | Can recipes/methods/protocols be versioned within the application? | ☐ Yes: Score = 1 ☐ No: Score = 3 ☐ N/A: Score = N/A |  |  |  |  |
| Average Control Element Score: |  |  |  |  |  | Report average score, rounded up/down to the nearest whole number. |
| Patient Impact (overall): |  |  |  |  |  | What is the impact of any control element with a score of 3 or greater? |
| Detectability (overall): |  |  |  |  |  | What is the detectability of any control element with a score of 3 or greater? |
| Overall Risk Rating = Average Control Element Score x Patient Impact x Detectability Score: |  |  |  |  |  |  |
| Additional Comments: |  |  |  |  |  |  |
| Table completed by (Initial and Date): |  |  |  |  |  |  |

# System Governance

| Table 11 – System Governance |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Item | Control Element |  | Meets Requirement (Yes / No or N/A) | Score |  | Outline mitigation strategy or risk-based justification for any control element that scores 5 or comment for clarification purposes. |
| System Validation |  |  |  |  |  |  |
|  | Have 21 CFR Part 11 and Annex 11 requirements been included in validation of the system? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
|  | Is periodic review of the system defined in a procedure? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
|  | Does a data process map exist in the validation document? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
|  | Does a data lifecycle diagram exist? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
| System Audit Trail Review |  |  |  |  |  |  |
|  | Is there a procedure that describes the system audit trail review and a defined frequency? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
|  | Does the procedure include review of configuration changes, data modification/movement, changes to calculations and user access; and are specific flags/symbols/acronyms defined? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
|  | Does the system audit trail review process include Quality oversight? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
| User Administration |  |  |  |  |  |  |
|  | Is there a procedure that describes the defined user roles and defines assignment of user roles? |  | ☐ Yes: Score = 1 ☐ No: Score = 3 ☐ N/A: Score = N/A |  |  |  |
|  | Is there a procedure that defines how users are added to the system? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
|  | Are user accounts approved by area management or system owner? |  | ☐ Yes: Score = 1 ☐ No: Score = 3 ☐ N/A: Score = N/A |  |  |  |
|  | Is there a procedure for the routine review of user accounts and roles for the system? |  | ☐ Yes: Score = 1 ☐ No: Score = 3 ☐ N/A: Score = N/A |  |  |  |
|  | Is there a procedure that defines the roles and activities of the system administrator? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
| Backup |  |  |  |  |  |  |
|  | Is there a procedure that defines the frequency of verification of the restore process and accessibility of backup and archived data? |  | ☐ Yes: Score = 1 ☐ No: Score = 3 ☐ N/A: Score = N/A |  |  |  |
|  | Is the backup process included in validation? |  | ☐ Yes: Score = 1 ☐ No: Score = 3 ☐ N/A: Score = N/A |  |  |  |
|  | Is the restore process included in validation? |  | ☐ Yes: Score = 1 ☐ No: Score = 3 ☐ N/A: Score = N/A |  |  |  |
| Data Review |  |  |  |  |  |  |
|  | Is there a procedure that defines how to review reported data from the system? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
|  | Does the review procedure include the review of audit trails that capture activities during the data acquisition and reporting process? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
|  | If data can be stored in multiple folders by the operator does the data review procedure include checks to determine if data was stored in folders/locations other than the designated locations? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
|  | Are critical equipment/system parameters verified during review? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
|  | Is data review performed using a paper printout or PDF copy of the results only? |  | ☐ Yes: Score = 5 ☐ No: Score = 1 ☐ N/A: Score = N/A |  |  |  |
|  | If paper printouts or PDF files are used to support data review, is there verification that no other original raw data was not reported? |  | ☐ Yes: Score = 1 ☐ No: Score = 3 ☐ N/A: Score = N/A |  |  |  |
|  | Is data review done using the electronic records? |  | ☐ Yes: Score = 1 ☐ No: Score = 3 ☐ N/A: Score = N/A |  |  |  |
|  | Is the audit trail related to data generation reviewed? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
|  | Is the audit trail review process and frequency outlined in a procedure? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
|  | Is the data electronically signed by the reviewer as having been reviewed? |  | ☐ Yes: Score = 1 ☐ No: Score = 3 ☐ N/A: Score = N/A |  |  |  |
|  | Has the type of data (static vs. dynamic) been defined/determined? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
| Control of Methods/Recipes |  |  |  |  |  |  |
|  | Is there a procedure that outlines how recipes/methods are managed for the system? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
|  | Do method changes outside of validated process parameters or not defined in the validation require QA approval and is this process described in a SOP? |  | ☐ Yes: Score = 1 ☐ No: Score = 5 ☐ N/A: Score = N/A |  |  |  |
|  | Are method templates/recipes reviewed on a defined basis for accuracy? |  | ☐ Yes: Score = 1 ☐ No: Score = 3 ☐ N/A: Score = N/A |  |  |  |
| Average Control Element Score: |  |  |  |  |  | Report average score, rounded up/down to the nearest whole number. |
| Patient Impact (overall): |  |  |  |  |  | What is the impact of any control element with a score of 3 or greater? |
| Detectability (overall): |  |  |  |  |  | What is the detectability of any control element with a score of 3 or greater? |
| Overall Risk Rating = Average Control Element Score x Patient Impact x Detectability Score: |  |  |  |  |  |  |
| Additional Comments: |  |  |  |  |  |  |
| Table completed by (Initial and Date): |  |  |  |  |  |  |

| Additional Comments: |  |
| --- | --- |
| Table completed by (Initial and Date): |  |

# System Data Integrity Risk Summary:

Summarize the reported scoring from each of the 10 categories in the table below. This represents the data integrity risk profile for the system. The overall Data Integrity Risk Score is an average score of all 10 categories.

| System Data Integrity Risk Profile |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| OS Security | Software Security | Data Back-Up | Audit Trail Criteria |  | Data Folder Security |  | Control of Date / Time | User Levels and Privileges | Password Control | Control of Methods | System Governance |
| Overall Data Integrity Risk Score: |  |  |  |  |  |  |  |  |  |  |  |
| Reported by (initial & date): |  |  |  |  |  |  |  |  |  |  |  |

Risk Remediation

| Final Risk Rating |  |  |
| --- | --- | --- |
| Risk Rating | Action | Mitigation |
| >25 | High Risk-Potential Impact to Patient Safety or Product Quality | Mandatory |
| 12-25 | Moderate Risk-No Impact to Patient Safety or Product Quality but Potential Regulatory Risk | Recommended |
| <12 | Negligible DI Risk | Not Required |

- Based on the above action table, determine if risk remediation strategies are required.

- Strategies to mitigate risk will be determined by the Process/System Owner and/or SME in consultation with Quality Assurance (Data Integrity Officer/ Manager), and Technical Operations as appropriate, to ensure risk reduction.

- In the case of long-term risk remediation actions, risk reducing short-term actions shall be implemented to reduce risk and provide an acceptable level of governance until the long-term remediation actions are completed.

- Relevant site procedures (e.g., change control, validation policy) will outline the scope of additional testing through the change management process.

- Reassessment of the system may be completed following the completion of remediation activities. The reassessment may be done at any time during the remediation process to document the impact of the remediation actions.

- Once final remediation is complete, a reassessment of the equipment/system will be completed to demonstrate that the risk rating has been mitigated by the remediation actions taken.